Joint Stock Company “LENTA”

# MEDICAL BANDAGES

Technical specifications

ТУ BY 700002794.008-2008

2008

### ТУ BY 700002794.008-2008 с.2

Technical specifications of medical bandages intend for medical treatment, fractures fixation, homostassis, avoid infection and etc.

# 1 Technical requirements

1.1 Our medical bandages conform to these Technical specifications, correspond to Government Standards (ГОСТ 15.007) and produced in accordance with confirmed technical documents.

1.2 Medical bandages of JSC “LENTA”:

- Fixing knitted medical bandages;

- Fixing woven medical bandages;

- Elastic knitted medical bandages;

- Elastic woven medical bandages;

- Tubular medical bandages;

- Fixing tubular elastic medical bandages;

- Medical bandages for gauze turunda.

1.3 We produce medical bandages of row material approved by Belarusian Health Authorities.

Our bandages are produced in different width, structure and application.

1 Physicochemical property

| Applications | Value |
| --- | --- |
| Water reaction | neutral |
| Wettability, с., | Maximum 10 |
| Capillarity, sm/h | Minimum 7 |
| Ash content, % | Maximum 0,3 |

| 2 Physical mechanical property |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Item | Width, mm | Length, m | Surface density, g/m², minimum | Breaking weight / Н, minimum | Material |
| 1. Fixing knitted medical bandages | 50±5 70±5 100±5 120±5 140±10 | 5; 7; 10 | 34 | 60 | Bleached cotton |
| 2. Fixing woven medical bandages | 50±5 70±5 100±5 | 5; 7; 10 | 34 | 60 | Bleached cotton |
| 3. Medical bandages for gauze turunda | 10±2 | 10; 20 | 50 | 15 | Bleached cotton |
|  | 20±2 |  |  |  |  |

| 3 Physical mechanical property |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Width, mm | Length, m | Surface density, g/m² | Breaking weight / Н, minimum | Stretching, % | Material |
| 1. Elastic knitted medical bandages | 95±5 | 2; 2,5; 3 | Minimum 250 Maximum 500 | 80 | Minimum 100, Maximum 350 | Polyamide Polyester Latex |
|  |  |  |  |  |  | Polyester Viscose Flax Polyamide Latex |
|  |  |  |  |  |  | Polyamide Cotton Latex |
| 2. Elastic woven medical bandages | 40±3 60±5 70±5 80±5 100±5 | 2; 2,5; 3 | Minimum 40, Maximum 150 | 80 | Minimum 50, Maximum 250 | Polyamide Viscose |
|  |  |  |  |  |  | Polyamide Cotton Viscose |

Note: Elastic woven and knitted medical bandages, fixing woven and knitted medical bandages are produced in different width and length according to customer requirements.

1.6 Allowance variation of Surface density:

- for fixing medical bandages — maximum ±5%;

- for elastic and tubular medical bandages – maximum ±8%.

1.7 Allowance variation of medical bandages length – maximum 2%.

1.8 It is not allowed in medical bandages any defects: rents and gaps, stains, contaminated yarns, missing weft – maximum 4 ends.

1.9 Allowable defects:

- kinking edges up to 1mm,

- missing 1 thread,

- threads thickening,

- nick up to 5mm,

- warp and weft threads separation,

- flexion up to 3cm for 1 meter bandage (except tubular bandages),

- cross stripes.

1.10 Our medical bandages safety and unirritate to skin.

New bandaging materials conform to EU Standards and have more sanitary and functional properties comparing with existent analogues.

| №4 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Size | Length, m | Surface density, g/m² | Breaking weight / Н, minimum | Stretching, % | Material |
| 1. Tubular medical bandages | 0-5 | 15 | Minimum 50 Maximum 300 | 80 | Non specified | Viscose |
|  |  |  |  |  |  | Cotton |
|  | 6-8 | 6 |  |  |  |  |
|  |  |  |  |  |  | Cotton Viscose |
| 2. Fixing tubular elastic medical bandages | 0-6 | 1; 2 | Minimum 15 Maximum 450 | 80 | Minimum 200 | Polyamide Lycra |
|  |  |  |  |  |  | Polyamide Latex |

A consignment of goods delivered with Quality certificate.

Every three years Sanitation Service checked our production.

Warranty self life of medical bandages 5 years.

Medical bandages of JSC “LENTA” were approbated by Byelorussian medical institutions. It was proved that our medical bandages reduce negative consequences of surgical treatable, anticipate the process of intention healing, burns, cuts, cracks, reduction of costs medical treatment and give us the possibility to appreciate the comfortable treatment.